## Slide 1: Uncertainties in Energy Rating (ER) and Module Performance Ratio (MPR)
B.V. Mihaylov, M. Bliss*, R. Gottschalg
Centre for Renewable Energy Systems Technology (CREST), Loughborough University
*Email: M.Bliss@lboro.ac.uk ; Tel.: +44 1509 635327
7th Energy Rating and Module Performance Modelling Workshop
30/03/2017

## Slide 2: Outline
  Introduction
  Uncertainty
    Methodology
    Sources and Correlations
    ER and MPR estimates
  Conclusions

## Slide 3: Introduction
For ER and MPR to be adopted and meaningful, the uncertainty analyses need to be consistent:
  worst-case assumptions may result in overestimation
  underestimating correlations may underestimate
Focus on uncertainty in performance matrix
  Methodology detailed
  Example given for CREST
Notes: Device measurements are used in combination with standard climatic datasets for calculating the ER and the MPR.
ER and MPR uncertainties are dominated by the performance matrix and Faiman coefficient measurement uncertainties.

## Slide 4: ER and MPR uncertainty methodology

## Slide 5: ER and MPR uncertainty methodology

## Slide 6: Uncertainty sources and correlation estimates
systematic uncertainty to all measurements
systematic uncertainty at same temperature
systematic uncertainty at same irradiance
is the random uncertainty
 | 15°C | 25°C | 50°C | 75°C
 | Uncertainty at k = 1 |  |  | 
1100 W/m2 | Na |  |  | 
1000 W/m2 |  |  |  | 
800 W/m2 |  |  |  | 
600 W/m2 |  |  |  | 
400 W/m2 |  |  |  | Na
200 W/m2 |  |  | Na | Na
100 W/m2 |  |  | Na | Na
 | 15°C | 25°C | 50°C | 75°C
1100 W/m2 | Na |  |  | 
1000 W/m2 |  |  |  | 
800 W/m2 |  |  |  | 
600 W/m2 |  |  |  | 
400 W/m2 |  |  |  | Na
200 W/m2 |  |  | Na | Na
100 W/m2 |  |  | Na | Na

## Slide 7: Uncertainty sources systematic to all measurements
uncertainty in reference cell calibration
uncertainty due to possible drift since the last calibration
uncertainty due to non-uniformity of the simulator (no filter)
uncertainty in voltage due to uncertainty in irradiance
is the uncertainty due to spectral mismatch at STC.
 |  |  |  |  |  | Total
 | 0.46% | 0.22% | 0.75% | 0.04% | 0.8% | 1.21%
 |  |  |  |  |  | Total
 | 0.46% | 0.22% | 0.75% | 0.04% | 0.8% | 1.21%

## Slide 8: Uncertainty sources systematic at the same temperature
is the systematic component of effective temperature uncertainty
is a typical coefficient for the given technology
is the uncertainty in MMF due to the change in SR of the DUT
is different at different set temperatures.
 |  |  |  | Total
 | 0.22°C | 0.45%/°C | 0.1% | 0.14%
 | 0.22°C | 0.45%/°C | 0% | 0.1%
 | 0.22°C | 0.45%/°C | 0.3% | 0.32%
 | 0.22°C | 0.45%/°C | 0.5% | 0.515%
 |  |  |  | Total
 | 0.22°C | 0.45%/°C | 0.1% | 0.14%
 | 0.22°C | 0.45%/°C | 0% | 0.1%
 | 0.22°C | 0.45%/°C | 0.3% | 0.32%
 | 0.22°C | 0.45%/°C | 0.5% | 0.515%

## Slide 9: Uncertainty sources systematic at the same irradiance
uncertainty due to possible non-linearity of the RC
uncertainty due to alignment
uncertainty due to orientation of the module
non-uniformity uncertainty introduced by the attenuation masks
is the uncertainty in MMF that is systematic with irradiance
is the uncertainty due to connecting the module
 |  |  |  |  |  |  | Total
 | 0.4% | 0.05% | 0.2% | 0.69% | 0.2% | 0.35% | 0.92%
 | 0.4% | 0.05% | 0.2% | 0.69% | 0.2% | 0.35% | 0.92%
 | 0.3% | 0.05% | 0.2% | 0.58% | 0.2% | 0.35% | 0.79%
 | 0.2% | 0.05% | 0.2% | 0.58% | 0.2% | 0.35% | 0.76%
 | 0.1% | 0.05% | 0.2% | 0.58% | 0.2% | 0.35% | 0.74%
 | 0% | 0.05% | 0.2% | 0% | 0% | 0.35% | 0.41%
 | 0.1% | 0.05% | 0.2% | 0% | 0.4% | 0.35% | 0.58%
 |  |  |  |  |  |  | Total
 | 0.4% | 0.05% | 0.2% | 0.69% | 0.2% | 0.35% | 0.92%
 | 0.4% | 0.05% | 0.2% | 0.69% | 0.2% | 0.35% | 0.92%
 | 0.3% | 0.05% | 0.2% | 0.58% | 0.2% | 0.35% | 0.79%
 | 0.2% | 0.05% | 0.2% | 0.58% | 0.2% | 0.35% | 0.76%
 | 0.1% | 0.05% | 0.2% | 0.58% | 0.2% | 0.35% | 0.74%
 | 0% | 0.05% | 0.2% | 0% | 0% | 0.35% | 0.41%
 | 0.1% | 0.05% | 0.2% | 0% | 0.4% | 0.35% | 0.58%

## Slide 10: Uncertainty sources random to all measurements
uncertainty due to the DAQ of the RC
uncertainty due to the DAQ of the DUT current
uncertainty due to the DAQ of the DUT voltage
random component of the MMF uncertainty
the uncertainty due to the fit of the MPP
repeatability uncertainty component,
random component of effective temperature uncertainty
typical TC.
 |  |  |  |  |  |  |  | Total
 | 0.12% | 0.12% | 0.12% | 0.1% | 0.06% | 0.35% | 0.26% | 0.50%
 |  |  |  |  |  |  |  | Total
 | 0.12% | 0.12% | 0.12% | 0.1% | 0.06% | 0.35% | 0.26% | 0.50%

## Slide 11: ER and MPR uncertainty methodology

## Slide 12: Covariance / Correlation between measurements
The covariance of two random variables X and Y, describes the level of linear dependence between them
Correlation is the normalized covariance
Covariance matrix is populated with the appropriate value depend on temperature and irradiance:

## Slide 13: Correlation matrix of 22 measurements

## Slide 14: ER and MPR uncertainty methodology

## Slide 15: Sample Performance matrices
Irradiance in W/m2
Temp in ºC
Max Power

## Slide 16: ER and MPR uncertainty methodology

## Slide 17: ER uncertainty with calculated correlations
ER uncertainty is smaller than that at STC
Strong correlation leads to lower MPR uncertainty
Climatic Data set | ER, kWh | ,% |  |  | ,%
Central UK , CREST | 271.08 | 1.258 | 0.858 | 0.964 | 0.711
Climatic Data set | ER, kWh |  |  |  | 
Central UK , CREST | 271.08 | 1.258 | 0.858 | 0.964 | 0.711

## Slide 18: ER uncertainty if all measurements are fully correlated
Overestimation of ER uncertainty
MPR uncertainty unrealistically small
Climatic Data set | ER, kWh | ,% |  |  | ,%
Central UK , CREST | 271.09 | 1.507 | 1.000 | 0.964 | 0.145
Climatic Data set | ER, kWh |  |  |  | 
Central UK , CREST | 271.09 | 1.507 | 1.000 | 0.964 | 0.145

## Slide 19: ER uncertainty if all measurements are independent
Underestimation of ER uncertainty
MPR uncertainty very high
Climatic Data set | ER, kWh | ,% |  |  | ,%
Central UK , CREST | 271.08 | 0.455 | 0.158 | 0.964 | 1.393
Climatic Data set | ER, kWh |  |  |  | 
Central UK , CREST | 271.08 | 0.455 | 0.158 | 0.964 | 1.393

## Slide 20: … up next …
  Introduction
  Uncertainty
    Methodology
    Sources and Correlations
    ER and MPR estimates
  Conclusions

## Slide 21: Conclusions
Measurements would have a significant commercial value
Need a consistent and robust uncertainty analysis
A framework for calculating the uncertainty is provided
Estimating the measurement correlation is key in correctly calculating the overall uncertainty
Alongside a procedure for calculating ER and MPR, there should be a procedure / guideline for propagating the uncertainty

## Slide 22: …Thank You…

## Slide 23: Difference in variance between extrapolated and interpolated conditions