## Slide 1
Uncertainty of DSR measurements according to approximations defined in the IEC 60904-8 standard
  K. Bothe, D. Hinken and C. Schinke
  Calibration and Test Center Solar Cells
  Institute for Solar Energy Research
  SUPSI-Workshop, Photoclass Project (4/2017)

## Slide 2: DSR system at ISFH CalTeC
Grating monochromator: 280 to 1200nm in 10nm steps
48 bias lamps, bias current up to 14A for large-area solar cells
Three transimpedance amplifiers (small: 250mA, large: 14A, Vmon)
Two light fields: 50x50mm² and 160x160mm²
Motorized axis for reference and sample cells
Sample Temperatures from 20 to 40°C (determination of TC)
ISO 17025 accredited by DAkkS since 2016

## Slide 3: Measurement procedure
Calibration of monochromatic light and bias light (Ebias) using a WPVS reference solar cell
DSR measurement of device under test at various (usually 8) bias levels: 10, 100, 200, 400, 600, 800, 1000 and 1100 W/m²
Integration over Ebias
Calculation of relative sstc-curves, sstc.rel(l) and mismatch correction factor
Determination of ISTC at sun simulator
Scaling of DSR curves and sstc.rel(l) using ISTC
Difference to PTB approach (previous talk by I. Kröger): Integration is not carried out over Ibias since only relative (unscaled) DSR values are measured

## Slide 4: Measurement uncertainty
Monte-Carlo uncertainty analysis with 12 uncertainty components:
fdist: Height-difference of reference and DUT
fwlshift: Deviation in wavelength of monochromatic light
fbandwidth: Bandwidth of monochromatic light
fTRef: Temperature difference to 25°C of reference
fTDUT: Temperature difference to 25°C of DUT
fcellinhom: Impact of light inhomogeneity on cells with current collection inhomogeneity
frepRef: Reproducibility of measurement of reference
frepDUT: Reproducibility of measurement of DUT
fscale: Uncertainty of Isc from IV measurement
fnonlin: Non-linearity of transimpedance amplifier
fref: Uncertainty of primary normal
fhom: Reproducibility of inhomogeneity correction

## Slide 5: Comparison to PTB: WPVS reference solar cell
WPVS reference solar cell
Fixed bias intensity/current
Curves of PTB (black) and ISFH-CalTeC (red)
Enavg = 0.1

## Slide 6: Solar Cell Calibration Standards
F. D‘Amore, Solar standards and certification, www.med-desire.eu, 2015

## Slide 7: IEC 60904-8 Ed. 3.0
F. D‘Amore, Solar standards and certification, www.med-desire.eu, 2015

## Slide 8: IEC 60904-8 Ed. 3.0
F. D‘Amore, Solar standards and certification, www.med-desire.eu, 2015
Required for spectral mismatch correction

## Slide 9: IEC 60904-8 Ed. 3.0  complete DSR procedure
Definition of the requirements for the measurement of the spectral responsivity of linear and non-linear photovoltaic devices:
  For highest accuracy, the differential spectral responsivity has to be measured under at least 5 different bias light irradiances resulting in short circuit currents between 5% and 110% of the short circuit current under standard test conditions . The spectral responsivity is calculated by integrating over .
   complete differential spectral responsivity (DSR) procedure

## Slide 10: IEC 60904-8 Ed. 3.0  simplifications
Simplifications aiming at determining one or more appropriate bias irradiances at which the measured differential spectral responsivity best approximates the spectral responsivity
  bias ramps at 3 to 5 wavelength with step width of 200nm increasing the bias light irradiance in 3 to 5 steps corresponding to between 5% and 110% of  multicolor bias ramps
  use of white light instead of monochromatic light  white bias ramp
  bias irradiance resulting in a bias current between 30% to 40% of  30% to 40% bias
  bias irradiance resulting in a bias current of 10% of if linearity is proven by showing that the differential spectral responsivity does not change by more than 2% when measuring at bias light intensities corresponding to 5% and 15% (not considered: only non-linear cells analyzed here)

## Slide 11: IEC 60904-8 Ed. 3.0  simplifications
Simplifications aiming at determining one or more appropriate bias irradiances at which the measured differential spectral responsivity best approximates the spectral responsivity
  bias ramps at 3 to 5 wavelength with step width of 200nm increasing the bias light irradiance in 3 to 5 steps corresponding to between 5% and 110% of  multicolor bias ramps
  use of white light instead of monochromatic light  white bias ramp
  bias irradiance resulting in a bias current between 30% to 40% of  30% to 40% bias
  bias irradiance resulting in a bias current of 10% of if linearity is proven by showing that the differential spectral responsivity does not change by more than 2% when measuring at bias light intensities corresponding to 5% and 15% (not considered: only non-linear cells analyzed here)
9-25 (+1)
3-5 (+1)
1
3
no. of measurements

## Slide 12: Deviations of simplifications compared to complete DSR procedure
Simulation of the DSR of a non-linear c-Si solar cell and analysis according to the complete DSR procedure as well as simplifications 1 to 3
Measurement of the DSR of a non-linear c-Si solar cell and analysis according to the complete DSR procedure as well as simplifications 1 to 3

## Slide 13: Simulation approach
p-type Cz Si
τSRH,n0 = 80 µs
τSRH,p0 = 800 µs
J0r,c = 790 fA/cm²
J0e = 59 fA/cm²
SiO
Sn = 1.22×104 cm/s
Sp = 5.92 cm/s
FEM simulation of a PERC c-Si solar cell using SENTAURUS DEVICE
Silicon dioxide dielectric layer at the rear side with very high interface defect density of 3×1010 cm-2
curves show high non-linearity
Bias ramps at different wavelengths yield bias intensity setpoints E0 from 286 to 317 W/m²

## Slide 14: Impact of bias ramp wavelength and bias irradiance
How much do the simplifications deviate from the complete DSR method?
Simplification 1 (Multicolor bias ramps): Deviations below -1.3%
Simplification 2 (White bias ramp): Deviations below 4.6%
Simplification 3 (30% bias): Deviations below 3.9%

## Slide 15: Measurement
Si3N4
p-type Cz Si
p-type Cz Si without AlOx but with SiN
curves show high non-linearity
Bias ramps at different wavelengths yield bias intensity setpoints from 287 to 301 W/m² (Simulation: 286 – 317 W/m²)

## Slide 16: Impact of bias ramp wavelength and bias irradiance
How much do the simplifications deviate from the complete DSR method?
Simplification 1 (Multicolor bias ramps): Deviations below -0.2%
Simplification 2 (White bias ramp): Deviations below -1%
Simplification 3 (30% bias): Deviations below -1%

## Slide 17: Summary
Analysis of non-linear c-Si solar cell (simulation and measurement).
Deviations below 5% were determined from solar cell device simulations for all approximations.
Simplification 1 (Multicolor-biasramps) was the most robust approach (deviations below 1.3%).
Simplification 2 (White-biasramp) showed deviations below 4.6%.
Simplification 3 (30% bias) showed deviations below 3.9%.
For non-linear solar cells: Use the complete DSR procedure if possible.
If a simplification is required, use the multicolor-biasramps approach if possible.
Thank you for your attention!

## Slide 18: Bias light intensities