## Slide 1: Performance of PV modules under different irradiances and temperatures
Robert Kenny

## Slide 2
Why the need for Energy Rating
      Peak power (Wp) at Standard Test Conditions (STC) is currently the standard performance guide for modules
      End users need kWh NOT Wp!
Energy Rating Procedure
      The performance is measured over a range of irradiances and temperatures to simulate the conditions that will be experienced outdoors
      Outdoor verification: The module is placed outdoors on the Energy Rating (ENRA) fixed rack and continuously monitored for up to one year

## Slide 3: Energy Rating requires Power Rating at different conditions
STC measurement:
  25°C
  1000W/m2
  AM1.5G
  Where:
  g=[100,200,400,600,800,1000,1100]W/m2
  k=[15,25,50,75]°C
  POWER RATING
  IEC 61853-1

## Slide 4
Characterising PV modules
IEC 61853-1 specifies:
      Ranges of Irradiance and Temperature
      Different measurement methods
      Interpolation methods
      But:
      Different methods cannot always meet the full ranges
      Different measurement systems have particular limitations, e.g. ranges and uncertainties
      Not all measurement systems are suitable for all module technologies (e.g. steady-state Vs pulsed simulators)

## Slide 5: IEC 61853 Part 1
POWER RATING
(FLAT PLATE MODULES)
61853-1 PHOTOVOLTAIC (PV) MODULE PERFORMANCE TESTING AND ENERGY RATING Part 1: Irradiance and temperature performance measurements and Power Rating
Power versus Irradiance and Temperature
IRR | Spectrum | Module Temperature |  |  | 
W-m-2 |  | 15C | 25C | 50C | 75C
1100 | AM1.5 | NA |  |  | 
1000 | AM1.5 |  |  |  | 
800 | AM1.5 |  |  |  | 
600 | AM1.5 |  |  |  | 
400 | AM1.5 |  |  |  | NA
200 | AM1.5 |  |  |  | NA
100 | AM1.5 |  |  | NA | NA

## Slide 6
IEC 61853-1 methods
      ‘Simplified’ procedure (linear modules)
        IEC 60904-10
      Natural sunlight with tracker
        Mesh filters/angle of incidence
      Natural sunlight without tracker
      Solar simulator
        Distance/angle of incidence/mesh filters (calibrated or uncalibrated)/decaying flash

## Slide 7
Measurement systems at ESTI
      Irradiance variation methods
      Simulators
      Pulsed: meshes/masks/flash decay
      Steady-state: meshes/# lamps/lamp voltage
      Natural sunlight with tracker
      Mesh filters
      Natural sunlight without tracker
      ‘The weather’

## Slide 8: IEC 61853-1 Indoor method
  g=[100,200,400,600,800,1000]W/m2
  k=[25,35,45,55,65]°C
  <<Procedure with solar simulator>>
  decay
  masks
  P3B
  heated box
  P2B
  varying G
  varying T
  none
  Time (ms)
  Irradiance
  P2B LAPSS pulse characteristic
  20
  0

## Slide 9: Mesh Filtering (uncalibrated)
  <<uncalibrated mesh filters>>
Mesh Assembling | Typical measured mean irradiance (W/m2) | Nominal corrected irradiance (W/m2)
A | ≈ 740 | 800
Bx | ≈ 635 | 600
B1·B2 | ≈ 405 | 400
A·B1·B2·Cx | ≈ 195 | 200
A·B1·B2·C1·C2 | ≈ 125 | 100
  Constraints:
  IEC 60891 within 30%
  IEC 61853-1 less than 100W/m2
900
1300
150

## Slide 10: Outdoor with tracker
  g=[100,200,400,600,800,1000]W/m2
  k=[25,35,45,55]°C
  <<under natural sunlight with tracker>>
  Temperature:
  Heated up by sun
  Cooled down with water
  Fine control using an opaque lid
  Irradiance:
  Mesh Filters
Mesh filters
D.U.T.
Aluminum structure
RefCell
Wooden box

## Slide 11: Outdoor field (ENRA)
  G high & T high
  G low & T low
  <<under natural sunlight without tracker>>
  BINNING TRESHOLDS:
  G within ± 2% the target irradiance
  T within ± 1°C the target temperature
  No Rs,
  No Spectral Corrections

## Slide 12: IV Curve Corrections
  IEC 60891
  Rs effect
  740  800 W/m2
  MMF correction
  IEC 60904-3,7
  NO Temperature corrections

## Slide 13: Results
  decay, mesh on P2B
  masks, mesh on P3B
  varying G,
  fixed T=25±0.5°C
  4 methods:
  Poly-Si
  CdTe
  Dev(Pmax) ≈ 5%
  Dev(Pmax) ≈ 3%
  @100 W/m2
  @100 W/m2

## Slide 14: Results (poly-Si)
Poly-Si | Mean Irradiance (W/m2) | Max Power (W)
1000 (W/m2) |  | 
PASAN STC1 | 998.5 | 54.2
PASAN STC2 | 995.2 | 54.1
PASAN IIIB STC1 | 1006.4 | 54.4
PASAN IIIB STC2 | 1008.7 | 54.3
Mean Power ± % Std Dev. | 54.25 W ± 0.27% | 
800 (W/m2) |  | 
PASAN DECAY | 742.5 | 43.4
PASAN MESH | 744.8 | 43.3
PASAN IIIB MASK | 704.3 | 43.7
PASAN IIIB MESH | 807.3 | 43.6
Mean Power ± % Std Dev. | 43.51 W ± 0.41% | 
600 (W/m2) |  | 
PASAN DECAY | 631.2 | 32.5
PASAN MESH | 631.9 | 32.4
PASAN IIIB MASK | 704.2 | 32.7
PASAN IIIB MESH | 602.4 | 32.6
Mean Power ± % Std Dev. | 32.54 W ± 0.40% | 
400 (W/m2) |  | 
PASAN DECAY | 405.1 | 21.3
PASAN MESH | 403.3 | 21.3
PASAN IIIB MASK | 402.3 | 21.5
PASAN IIIB MESH | 403.2 | 21.4
Mean Power ± % Std Dev. | 21.38 W ± 0.33% | 
200 (W/m2) |  | 
PASAN DECAY | 191.8 | 10.1
PASAN MESH | 189.8 | 10.0
PASAN IIIB MASK | 201.0 | 10.2
PASAN IIIB MESH | 197.8 | 10.5
Mean Power ± % Std Dev. | 10.20 W ± 2.42% | 
100 (W/m2) |  | 
PASAN DECAY | 126.7 | 4.7
PASAN MESH | 125.7 | 4.3
PASAN IIIB MASK | 101.0 | 4.7
PASAN IIIB MESH | 113.9 | 4.4
Mean Power ± % Std Dev. | 4.52 W ± 4.97% | 

## Slide 15: Results (CdTe)
CdTe | Mean Irradiance (W/m2) | Max Power (W)
1000 (W/m2) |  | 
PASAN STC | 995.0 | 68.0
PASAN IIIB STC1 | 974.2 | 67.5
PASAN IIIB STC2 | 974.3 | 67.6
Mean Power ± % Std Dev. | 67.68 W ± 0.39% | 
800 (W/m2) |  | 
PASAN DECAY | 798.2 | 54.8
PASAN MESH | 742.7 | 55.0
PASAN IIIB MASK | 779.7 | 54.8
PASAN IIIB MESH | 782.6 | 54.8
Mean Power ± % Std Dev. | 54.85 W ± 0.15% | 
600 (W/m2) |  | 
PASAN DECAY | 591.7 | 41.1
PASAN MESH | 628.6 | 41.4
PASAN IIIB MASK | 575.4 | 41.2
PASAN IIIB MESH | 579.2 | 41.2
Mean Power ± % Std Dev. | 41.25 W ± 0.32% | 
400 (W/m2) |  | 
PASAN DECAY | 400.0 | 27.0
PASAN MESH | 403.7 | 27.2
PASAN IIIB MASK | 388.2 | 27.2
PASAN IIIB MESH | 387.1 | 27.2
Mean Power ± % Std Dev. | 27.15 W ± 0.46% | 
200 (W/m2) |  | 
PASAN DECAY | 202.7 | 12.5
PASAN MESH | 189.9 | 12.7
PASAN IIIB MASK | 193.6 | 12.7
PASAN IIIB MESH | 188.6 | 12.4
Mean Power ± % Std Dev. | 12.58 W ± 1.11% | 
100 (W/m2) |  | 
PASAN DECAY | 101.2 | 5.3
PASAN MESH | 115.5 | 5.7
PASAN IIIB MASK | 96.9 | 5.5
PASAN IIIB MESH | 91.9 | 5.5
Mean Power ± % Std Dev. | 5.49 W ± 2.76% | 

## Slide 16: Non uniformity
  Low irradiance 100-200W/m2
  4 or 5 filters stacked together
  No spaces among filters
  Spatial uniformity issues
  Indoor Isc deviation > 10%
  Visible Moiré pattern:
800 W/m2
400 W/m2
100 W/m2
1 mesh
2 meshes
5 meshes

## Slide 17: Mesh Filtering - improved filters
  <<uncalibrated mesh filters>>
  1.2 m width

## Slide 18: Self-Reference method
  Isc of module measured with Pasan 3B as reference
  Linearity check (IEC 60904-10)
  MMF corrected
  Determination of Isc(T)
  Rs correction
  NO spectral mismatch (MMF=1)
  Further corrections:
  Envisaged by IEC 61853-1
  Advantages:
  ratio ADUT/Arefcell ≈ 100
  different geometry of the cells

## Slide 19: Pasan 3B improved meshes (c-Si)

## Slide 20: Pasan 3B improved meshes (CdTe)

## Slide 21: Comparing indoor and outdoor methods
  CdTe
  indoor
  (mesh & decay)
  outdoor tracker
  Average of 2
  data sets for
  each plotted
  surface
  Avg % difference:
  Poly-Si ~1%
  CdTe ~1%
  CIS ~2%

## Slide 22: Comparing outdoor methods
  Poly-Si
Pmax(W) | 25 | 35 | 45 | 55
100 | 5.03 | - | - | -
200 | 10.15 | 9.63 | - | -
400 | 21.74 | 20.62 | 18.74 | 17.69
600 | 32.66 | 31.74 | 30.09 | 27.06
800 | 42.57 | 41.27 | 39.72 | 36.31
1000 | - | - | 48.99 | 44.24
out-ENRA (%) | 25 | 35 | 45 | 55
100 | 2.28 | - | - | -
200 | 0.34 | 2.35 | - | -
400 | 1.35 | 1.92 | 3.16 | 3.57
600 | 0.81 | 2.94 | 2.39 | 3.05
800 | 1.71 | 0.55 | 2.33 | 2.54
1000 | - | - | 1.19 | 4.04
in-ENRA (%) | 25 | 35 | 45 | 55
100 | 5.41 | - | - | -
200 | 0.77 | 1.53 | - | -
400 | 0.08 | 0.49 | 4.31 | 5.03
600 | 0.12 | 1.60 | 1.00 | 4.17
800 | 2.35 | 0.91 | 0.15 | 3.58
1000 | - | - | 4.21 | 5.40
  Outdoor field
  AVG % dev:
  In-ENRA ≈ 2.5 %
  outrk-ENRA ≈ 2 %

## Slide 23: Comparing outdoor methods
  CIS
Pmax(W) | 25 | 35 | 45 | 55
100 | 2.28 | 2.16 | - | -
200 | 7.22 | 6.74 | - | -
400 | 19.82 | 19.23 | 18.61 | 17.70
600 | 33.93 | 33.11 | 32.19 | 31.07
800 | - | 47.98 | 46.34 | 44.11
1000 | - | 63.03 | 57.75 | 57.24
out-ENRA (%) | 25 | 35 | 45 | 55
100 | 13.05 | - | - | -
200 | 10.14 | 5.04 | - | -
400 | 5.64 | 6.44 | 6.03 | 3.80
600 | 3.79 | 4.60 | 4.26 | 2.64
800 | - | 4.77 | 3.89 | 1.96
1000 | - | 4.35 | 1.77 | 0.83
in-ENRA (%) | 25 | 35 | 45 | 55
100 | 7.98 | 3.94 | - | -
200 | 6.57 | 0.82 | - | -
400 | 4.48 | 3.76 | 1.79 | 1.40
600 | 3.09 | 3.04 | 2.41 | 1.38
800 | - | 4.06 | 3.52 | 1.52
1000 | - | 4.53 | 1.19 | 3.90
  Outdoor field
  AVG % dev:
  In-ENRA ≈ 3.3 %
  outrk-ENRA ≈ 4.8 %

## Slide 24: Comparing outdoor methods
  CdTe
Pmax(W) | 25 | 35 | 45 | 55
100 | 5.98 | 5.53 | - | -
200 | 13.21 | 12.96 | 12.91 | -
400 | 27.38 | 27.67 | 25.45 | 24.88
600 | 39.60 | 39.77 | 39.86 | 38.92
800 | 54.09 | 53.85 | 52.84 | 51.35
1000 | - | 65.58 | 64.92 | 64.37
out-ENRA (%) | 25 | 35 | 45 | 55
100 | 12.60 | 7.53 | - | -
200 | 6.76 | 7.83 | 8.90 | -
400 | 1.49 | 4.83 | 0.84 | 0.07
600 | 3.65 | 1.12 | 1.30 | 1.59
800 | 1.16 | 0.21 | 0.49 | 0.41
1000 | - | 1.75 | 0.27 | 1.40
in-ENRA (%) | 25 | 35 | 45 | 55
100 | 8.50 | 2.43 | - | -
200 | 4.80 | 4.79 | 6.73 | -
400 | 1.53 | 3.89 | 1.85 | 1.18
600 | 3.08 | 1.46 | 1.18 | 1.72
800 | 1.52 | 0.08 | 0.60 | 0.08
1000 | - | 2.74 | 1.34 | 0.46
  Outdoor field
  AVG % dev:
  In-ENRA ≈ 2.4 %
  outrk-ENRA ≈ 2.2 %

## Slide 25: What about PV systems?
Indoor-outdoor comparison of power matrices in the range of T and G considered:
c-Si modules: ±1.5% (after removing an offset of -2.9%)
CdTe modules: ±4.4% (after removing an offset of -4.1%)

## Slide 26: Results & Conclusions
  Mesh filtering technique is promising for outdoor measurements.
  IN-OUT agreement within 3%.
  Fixed-rack data, on average, deviation less than 5%.
  Self-Reference method improve results.
  Different technique appropriate for different technologies, but none ideal.
Issues
  Mesh filters availability/size.
  Non uniformity at low irradiance, especially for thin film modules.
  Outdoors cannot reach 1100Wm-2 in many sites.

## Slide 27
References
[1] Photovoltaic (PV) module performance testing and energy rating - Part1: Irradiance and temperature performance measurements and Power Rating, IEC61853-1, Edition 1.0, 2011-01
[2] Robert P. Kenny, Davide Viganó, Elena Salis, Matthew Norton, Harald Müllejans, Willem Zaaiman, ‘Power rating of photovoltaic modules including validation of procedures to implement IEC 61853-1 on solar simulators and under natural sunlight’, Prog. Photovolt: Res. Appl. 2013; 21:1384–1399
[3] Adrián A. Santamaría Lancia, Giorgio Bardizza, Harald Müllejans, ‘Assessment of uncalibrated light attenuation filters constructed from industrial woven wire meshes for use in photovoltaic research’, presented at this conference: paper number 5CV.2.2
References
[1] Robert P. Kenny, Davide Viganó, Elena Salis, Matthew Norton, Harald Müllejans, Willem Zaaiman, ‘Power rating of photovoltaic modules including validation of procedures to implement IEC 61853-1 on solar simulators and under natural sunlight’, Prog. Photovolt: Res. Appl. 2013; 21:1384–1399
[2] Adrián A. Santamaría Lancia, Giorgio Bardizza, Harald Müllejans, ‘Assessment of uncalibrated light attenuation filters constructed from industrial woven wire meshes for use in photovoltaic research’, presented at EUPVSEC conference